CALIBRATION CERTIFICATE
Date of calibration: 28.06.2018

Type: BDKG-02 detection unit

S/N: 428.002

Measurement limits:
BDKG-02: γ 0.1 µSv/h – 10 Sv/h;

Measurement error:
BDKG-02: ±15%

Operating conditions:
Air temperature	+21°C
Atmospheric pressure	100.5 kPa
Relative humidity	77%

The instrument is calibrated on AT-130 standard dosimetry facility, No. 015, error ±4% (Certificate of Compliance No.210-1099/15 as of 23.10.2015 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia); the instrument is calibrated on AT-110 standard dosimetry facility, No.013, error ±5% (Certificate of Compliance No.210-1098/15 as of 23.10.2015 issued by FGUP "D.I.Mendeleyev VNIIM", St.Petersburg, Russia) and standard gamma radiation 137Cs source of OSGI type with 9,0 kBq activity, s/n 340.05.2017.

Calibration data

BDKG-02 (γ)
Conventionally true value (CTV) of dose rate, / H_{0i}^{"*"}csup{%Ux2022} ( 10)	Calibration data, / H_{i}^{"*"}csup{%Ux2022} ( 10)
0.70 µSv/h	0.70 µSv/h
7.0 µSv/h	6.8 µSv/h
70.0 µSv/h	72.0 µSv/h
0.70 mSv/h	0.69 mSv/h
7.0 mSv/h	6.9 mSv/h
70.0 mSv/h	69.0 mSv/h
0.7 Sv/h	0.70 Sv/h
7.0 Sv/h	7.1 Sv/h

S/N	Нс calibration value, / nSv/h	Source activity, / kBq
428.002	921	1,03

Calibrated by:	V. Pisarenko	_____________
(Signature)

Technical control:	N. Kurbatova	_____________
(Signature)