CALIBRATION CERTIFICATE
AT2327 Alarm Dosimeter	Date of calibration: 04.12.2019.

SN: 497.100

Type: BDKG-204

Measurement limits:
BDKG-204: γ 0.05 µSv/h – 10 Sv/h;

Measurement error:
BDKG-204: ±20%

Operating conditions:
- Air temperature	+22,0 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	72,6 %
- Gamma radiation background	96 nSv/h

The instrument is calibrated on AT-130 standard dosimetry facility, No. 015, error ±4% (Certificate of Compliance No. 210-1427/18 as of 23.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia); the instrument is calibrated on AT-110 standard dosimetry facility, No.013, error ±5% (Certificate of Compliance No.210-1426/18 as of 22.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St.Petersburg, Russia)

Calibration data

BDKG-204 () sn: 497.101
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 µSv/h	263	200.7	96	0,073	0,069	0,068	0,070	0,00	5,50	±20
0,7 µSv/h	0НА	236	96	0,77	0,73	0,69	0,73	4,29	7,04
7 µSv/h	0НА	76	---	7,6	6,9	6,9	7,1	1,43	5,61
70 µSv/h	К/С Г3	257	---	73,4	71,4	65,3	70	0,00	5,50
0.7 mSv/h	К/С Г3	82	---	0,73	0,72	0,60	0,68	-2,86	6,41
7,0 mSv/h	043	357	---	7,1	7,3	7,2	7,2	2,86	5,50
70 mSv/h	043	114	---	65,2	74,1	70,7	70	0,00	4,40
0,7 Sv/h	163	235	---	0,66	0,71	0,66	0,68	-2,86	5,50
7 Sv/h	163	76	---	6,9	7,7	7,3	7,3	4,29	6,22

BDKG-204 () sn: 497.102
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 µSv/h	263	200.7	96	0,074	0,074	0,062	0,070	0,00	5,50	±20
0,7 µSv/h	0НА	236	96	0,69	0,67	0,65	0,67	-4,29	7,04
7 µSv/h	0НА	76	---	6,6	7,3	6,8	6,9	-1,43	5,61
70 µSv/h	К/С Г3	257	---	67,1	66,3	76,6	70	0,00	5,50
0.7 mSv/h	К/С Г3	82	---	0,64	0,62	0,74	0,67	-4,29	7,04
7,0 mSv/h	043	357	---	6,7	7,1	7,4	7,1	1,43	4,54
70 mSv/h	043	114	---	68,0	66,4	75,6	70	0,00	4,40
0,7 Sv/h	163	235	---	0,63	0,66	0,71	0,67	-4,29	6,22
7 Sv/h	163	76	---	6,9	7,2	6,8	7	0,00	4,40

BDKG-204 () sn: 497.103
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 µSv/h	263	200.7	96	0,075	0,075	0,061	0,070	0,00	5,50	±20
0,7 µSv/h	0НА	236	96	0,66	0,69	0,77	0,71	1,43	5,61
7 µSv/h	0НА	76	---	6,5	6,6	7,9	7	0,00	5,50
70 µSv/h	К/С Г3	257	---	72,7	71,6	65,7	70	0,00	5,50
0.7 mSv/h	К/С Г3	82	---	0,66	0,68	0,68	0,67	-4,29	7,04
7,0 mSv/h	043	357	---	7,0	7,0	7,4	7,1	1,43	4,54
70 mSv/h	043	114	---	73,3	66,6	70,1	70	0,00	4,40
0,7 Sv/h	163	235	---	0,66	0,64	0,73	0,68	-2,86	5,50
7 Sv/h	163	76	---	7,1	7,2	7,6	7,3	4,29	6,22

BDKG-204 () sn: 497.104
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 µSv/h	263	200.7	96	0,067	0,068	0,075	0,070	0,00	5,50	±20
0,7 µSv/h	0НА	236	96	0,75	0,74	0,64	0,71	1,43	5,61
7 µSv/h	0НА	76	---	7,3	6,9	6,9	7	0,00	5,50
70 µSv/h	К/С Г3	257	---	74,7	66,9	68,4	70	0,00	5,50
0.7 mSv/h	К/С Г3	82	---	0,73	0,68	0,63	0,68	-2,86	6,41
7,0 mSv/h	043	357	---	7,2	7,5	6,6	7,1	1,43	4,54
70 mSv/h	043	114	---	70,7	73,7	65,6	70	0,00	4,40
0,7 Sv/h	163	235	---	0,71	0,76	0,66	0,71	1,43	4,54
7 Sv/h	163	76	---	7,5	7,1	7,3	7,3	4,29	6,22

BDKG-204 () sn: 497.105
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 µSv/h	263	200.7	96	0,067	0,074	0,069	0,070	0,00	5,50	±20
0,7 µSv/h	0НА	236	96	0,75	0,72	0,66	0,71	1,43	5,61
7 µSv/h	0НА	76	---	7,0	7,4	6,6	7	0,00	5,50
70 µSv/h	К/С Г3	257	---	70,1	73,6	66,3	70	0,00	5,50
0.7 mSv/h	К/С Г3	82	---	0,70	0,69	0,62	0,67	-4,29	7,04
7,0 mSv/h	043	357	---	7,0	7,0	7,5	7,2	2,86	5,50
70 mSv/h	043	114	---	70,2	68,1	71,7	70	0,00	4,40
0,7 Sv/h	163	235	---	0,71	0,69	0,79	0,73	4,29	6,22
7 Sv/h	163	76	---	7,7	6,9	7,4	7,3	4,29	6,22

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)