CALIBRATION CERTIFICATE
Date of calibration: 18.07.2016.

Type: BDKG-11/1 detection unit

S/N: 352.002

Measurement limits:
BDKG-11/1: γ 0.01 µSv/h – 100 µSv/h;

Measurement error:
BDKG-11/1: ±20%

Operating conditions:
Air temperature	+21°C
Atmospheric pressure	99.5 kPa
Relative humidity	77%

The instrument is calibrated on AT-110 standard dosimetry facility, No.013, error ±5% (Certificate of Compliance No.210-1098/15 as of 23.10.2015 issued by FGUP "D.I.Mendeleyev VNIIM", St.Petersburg, Russia) and standard gamma radiation 137Cs source of OSGI-R type with 67,6 kBq activity, s/n 2910.

Calibration data

BDKG-11/1 (γ)

Conventionally true value (CTV) of dose rate, / H_{0i}^{"*"}csup{%Ux2022} ( 10)	Calibration data, / H_{i}^{"*"}csup{%Ux2022} ( 10)	Relative intrinsic error of a dosimeter, %
		At calibration, %	Limits of variation,%
0.07 µSv/h	0.071 µSv/h	1,4	±20
0.7 µSv/h	0.69 µSv/h	-1,4
7.0 µSv/h	7.0 µSv/h	0,0
70.0 µSv/h	73.0 µSv/h	4,3

S/N	Нс calibration value, / nSv/h (BDKG-11/1 in sealed container with protection)	Source activity, / kBq
352.002	1734	67,6

Calibrated by:	V. Pisarenko	_____________
(Signature)

Technical control:	N. Kurbatova	_____________
(Signature)

CALIBRATION CERTIFICATE
Date of calibration: 18.07.2016.

Type: BDKG-11/1 detection unit

S/N: 352.003

Measurement limits:
BDKG-11/1: γ 0.01 µSv/h – 100 µSv/h;

Measurement error:
BDKG-11/1: ±20%

Operating conditions:
Air temperature	+21°C
Atmospheric pressure	99.5 kPa
Relative humidity	77%

The instrument is calibrated on AT-110 standard dosimetry facility, No.013, error ±5% (Certificate of Compliance No.210-1098/15 as of 23.10.2015 issued by FGUP "D.I.Mendeleyev VNIIM", St.Petersburg, Russia) and standard gamma radiation 137Cs source of OSGI-R type with 67,6 kBq activity, s/n 2910.

Calibration data

BDKG-11/1 (γ)

Conventionally true value (CTV) of dose rate, / H_{0i}^{"*"}csup{%Ux2022} ( 10)	Calibration data, / H_{i}^{"*"}csup{%Ux2022} ( 10)	Relative intrinsic error of a dosimeter, %
		At calibration, %	Limits of variation,%
0.07 µSv/h	0.067 µSv/h	-4,3	±20
0.7 µSv/h	0.68 µSv/h	-2,8
7.0 µSv/h	7.0 µSv/h	0,0
70.0 µSv/h	71.5 µSv/h	2,1

S/N	Нс calibration value, / nSv/h (BDKG-11/1 in sealed container with protection)	Source activity, / kBq
352.003	1715	67,6

Calibrated by:	V. Pisarenko	_____________
(Signature)

Technical control:	N. Kurbatova	_____________
(Signature)