CALIBRATION CERTIFICATE

Type: Radiation monitor AT1117M	Date of calibration: 25.12.2019.

SN: 17777

Dose rate measurement limits of X-ray and gamma radiation:

BDKG-01: 0.1 Sv/h – 10.0 Sv/h;

03: ± 20 %БД
Measurement error:
BDKG-01: ± 20 %;

Operating conditions:
- Air temperature	+19,6 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	69,2 %
- Gamma radiation background	96 nSv/h

Calibration means:
- standard dosimetry facility AT-130, N 013, Cs-137, error 5%, the Certificate of Compliance N210/1426-2018 on 22.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia);

- standard dosimetry facility AT-130, N 015, Cs-137-137, error 4 %, the Certificate of Compliance N210/ 1427-2018 on 23.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia);

Calibration data
BDKG-01 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,7 Sv/h	0НА	236	96	0,70	0,66	0,74	0,70	0,00	5,50	±20
7 Sv/h	0НА	76	---	7,1	6,7	6,5	6,8	-2,86	6,33
70 Sv/h	К/С Г3	257	---	68,1	66,6	75,1	69,9	-0,14	5,50
0.7 mSv/h	К/С Г3	82	---	0,71	0,65	0,65	0,67	-4,29	7,24
7,0 mSv/h	043	357	---	7,6	7,1	6,6	7,1	1,43	4,67
70 mSv/h	043	114	---	70,7	71,9	67,7	70,1	0,14	4,40
0,7 Sv/h	163	235	---	0,71	0,73	0,63	0,69	-1,43	4,67
7 Sv/h	163	76	---	7,0	6,7	6,4	6,7	-4,29	6,45

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)