CALIBRATION CERTIFICATE
AT2327 Alarm Dosimeter	Date of calibration: 25.12.2019.

SN: 452.000

Type: BDKG-02

Measurement limits:
BDKG-02: γ 0.1 µSv/h – 10 Sv/h;

Measurement error:
BDKG-02: ±15%

Operating conditions:
- Air temperature	+20,7 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	72,4 %
- Gamma radiation background	96 nSv/h

The instrument is calibrated on AT-130 standard dosimetry facility, No. 015, error ±4% (Certificate of Compliance No. 210-1427/18 as of 23.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia); the instrument is calibrated on AT-110 standard dosimetry facility, No.013, error ±5% (Certificate of Compliance No.210-1426/18 as of 22.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St.Petersburg, Russia) and standard gamma radiation 137Cs source of OSGI type with 84,83 kBq activity, s/n №516.

Calibration data

BDKG-02 () sn: 452.002
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,7 Sv/h	0НА	236	96	0,67	0,68	0,69	0,68	-2,86	6,41	±15
7 Sv/h	0НА	76	---	6,8	7,0	7,5	7,1	1,43	5,61
70 Sv/h	К/С Г3	257	---	75,6	69,7	67,7	71,0	1,43	5,61
0.7 mSv/h	К/С Г3	82	---	0,66	0,67	0,75	0,69	-1,43	5,61
7,0 mSv/h	043	357	---	6,5	6,5	7,0	6,7	-4,29	6,22
70 mSv/h	043	114	---	72,5	73,2	67,3	71,0	1,43	4,54
0,7 Sv/h	163	235	---	0,69	0,69	0,59	0,66	-5,71	7,93
7 Sv/h	163	76	---	7,0	7,2	6,5	6,9	-1,43	4,54

S/N	Нс calibration value, / nSv/h	Source activity, / kBq

452.002	10320	84,83

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)