CALIBRATION CERTIFICATE
AT2327 Alarm Dosimeter	Date of calibration: 25.12.2019.

SN: 452.000

Type: BDKG-204

Measurement limits:
BDKG-204: γ 0.05 µSv/h – 10 Sv/h;

Measurement error:
BDKG-204: ±20%

Operating conditions:
- Air temperature	+21,5 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	74,8 %
- Gamma radiation background	99 nSv/h

The instrument is calibrated on AT-130 standard dosimetry facility, No. 015, error ±4% (Certificate of Compliance No. 210-1427/18 as of 23.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia); the instrument is calibrated on AT-110 standard dosimetry facility, No.013, error ±5% (Certificate of Compliance No.210-1426/18 as of 22.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St.Petersburg, Russia) .

Calibration data

BDKG-204 () sn: 452.003
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 Sv/h	263	200.7	99	0,072	0,074	0,062	0,069	-1,43	5,61	±20
0,7 Sv/h	0НА	236	99	0,69	0,73	0,68	0,70	0,00	5,50
7 Sv/h	0НА	76	---	6,8	6,9	7,7	7,1	1,43	5,61
70 Sv/h	К/С Г3	257	---	68,1	74,8	76,2	73,0	4,29	7,04
0.7 mSv/h	К/С Г3	82	---	0,67	0,71	0,75	0,71	1,43	5,61
7,0 mSv/h	043	357	---	6,8	6,8	5,9	6,5	-7,14	8,87
70 mSv/h	043	114	---	66,7	71,1	69,2	69,0	-1,43	4,54
0,7 Sv/h	163	235	---	0,71	0,69	0,76	0,72	2,86	5,50
7 Sv/h	163	76	---	7,1	7,5	6,7	7,1	1,43	4,54

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)