CALIBRATION CERTIFICATE
Type: BDKG-22	Date of calibration: 25.12.2019.

S/N: 123

Measurement limits:
BDKG-22: γ 0.1 µSv/h – 10 Sv/h;

Measurement error:
BDKG-22: ±15%

Operating conditions:
- Air temperature	+20,1 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	71,4 %
- Gamma radiation background	97 nSv/h

The instrument is calibrated on AT-130 standard dosimetry facility, No. 015, error ±4% (Certificate of Compliance No. 210-1427/18 as of 23.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia); the instrument is calibrated on AT-110 standard dosimetry facility, No.013, error ±5% (Certificate of Compliance No.210-1426/18 as of 22.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St.Petersburg, Russia)

Calibration data

BDKG-22 () sn: 123
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,7 Sv/h	0НА	236	97	0,71	0,74	0,66	0,70	0,00	5,50	±15
7 Sv/h	0НА	76	---	7,5	7,6	6,5	7,2	2,86	6,41
70 Sv/h	К/С Г3	257	---	75,4	72,1	71,5	73,0	4,29	7,04
0.7 mSv/h	К/С Г3	82	---	0,72	0,70	0,62	0,68	-2,86	6,41
7,0 mSv/h	043	357	---	6,5	6,1	7,1	6,6	-5,71	7,93
70 mSv/h	043	114	---	67,6	73,6	65,9	69,0	-1,43	4,54
0,7 Sv/h	163	235	---	0,74	0,72	0,67	0,71	1,43	4,54
7 Sv/h	163	76	---	7,2	6,9	7,8	7,3	4,29	6,22

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)