CALIBRATION CERTIFICATE
Type: BDKG-23/1	Date of calibration: 25.12.2019.

S/N:__125

Measurement limits:
BDKG-23/1: γ 0.1 µSv/h – 100 Sv/h;

Measurement error:
BDKG-23/1: ±15%

Operating conditions:
- Air temperature	+21,1 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	72,9 %
- Gamma radiation background	92 nSv/h

The instrument is calibrated on AT-130 standard dosimetry facility, No. 015, error ±4% (Certificate of Compliance No. 210-1427/18 as of 23.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia); the instrument is calibrated on AT-110 standard dosimetry facility, No.013, error ±5% (Certificate of Compliance No.210-1426/18 as of 22.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St.Petersburg, Russia)

Calibration data

BDKG-23/1 () sn: 125
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,7 Sv/h	0НА	236	92	0,68	0,70	0,70	0,69	-1,43	5,61	±15
7 Sv/h	0НА	76	---	6,5	7,1	7,0	6,9	-1,43	5,61
70 Sv/h	К/С Г3	257	---	71,6	74,6	72,8	73,0	4,29	7,04
0.7 mSv/h	К/С Г3	82	---	0,73	0,80	0,72	0,75	7,14	9,46
7,0 mSv/h	043	357	---	6,5	6,5	7,7	6,9	-1,43	4,54
70 mSv/h	043	114	---	69,1	67,7	64,2	67,0	-4,29	6,22
0,7 Sv/h	163	235	---	0,69	0,68	0,72	0,70	0,00	4,40
7 Sv/h	163	76	---	7,4	7,3	6,6	7,1	1,43	4,54
40 Sv/h	163	32.7	---	39,6	38,2	42,2	40,0	0,00	4,40

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)