CALIBRATION CERTIFICATE
Type: BDKG-23	Date of calibration: 25.12.2019.

S/N: 124

Measurement limits:
BDKG-23: γ 0.1 µGy /h – 100 Gy/h;

Measurement error:
BDKG-23: ±15%

Operating conditions:
- Air temperature	+20,5 °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	72,2 %
- Gamma radiation background	96 nGy/h

The instrument is calibrated on AT-130 standard dosimetry facility, No. 015, error ±4% (Certificate of Compliance No. 210-1427/18 as of 23.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia); the instrument is calibrated on AT-110 standard dosimetry facility, No.013, error ±5% (Certificate of Compliance No.210-1426/18 as of 22.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St.Petersburg, Russia)

Calibration data

BDKG-23 () sn: 124
Kerma rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Kerma rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nGy/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,7 Gy/h	УП4	304.7	96	0,67	0,72	0,73	0,71	1,43	5,61	±15
7 Gy /h	УП4	108.9	---	6,7	7,1	7,2	7,0	0,00	5,50
70 Gy /h	К/С Г3	233.5	---	64,9	65,5	70,6	67,0	-4,29	7,04
0.7 mGy /h	9TH	166.1	---	0,68	0,71	0,69	0,69	-1,43	5,61
7,0 mGy /h	9TH	53.6	---	7,2	6,9	6,3	6,8	-2,86	5,50
70 mGy /h	043	103.6	---	73,4	77,9	73,7	75,0	7,14	8,87
0,7 Gy/h	163	213.9	---	0,71	0,71	0,64	0,69	-1,43	4,54
7 Gy/h	163	68.8	---	7,1	6,8	7,7	7,2	2,86	5,50
40 Gy/h	163	30.7	---	43,8	40,2	39,0	41,0	2,50	5,50

Calibrated by:	V. Pisarenko
		(signature)
Technical control:	N. Kurbatova
		(signature)