CALIBRATION CERTIFICATE

Type: Radiation monitor AT1117M	Date of calibration: $date.

SN: $serial

Dose rate measurement limits of X-ray and gamma radiation:

PU detector: 1.0 Sv/h – 10.0 mSv/h;

PU detector: 1.0 Sv/h – 100.0 mSv/h;

PU-2 detector: 1.0 mSv/h – 100.0 mSv/h;

PU-2 detector: 1.0 mSv/h – 10.0 mSv/h;

PU-4 detector: 0.3 Sv/h – 0.1 Sv/h;

BDKG-01: 0.1 Sv/h – 10.0 Sv/h;

BDKG-03: 30.0 nSv/h – 0.3 mSv/h;

BDKG-04: 50.0 nSv/h – 10.0 Sv/h;

BDKG-05: 30.0 nSv/h – 0.3 mSv/h;

BDKG-11: 10 nSv/h – 0.1 mSv/h;

BDKG-17: 1.0 mSv/h – 100 Sv/h;

BDKG-24: 20 nSv/h – 1.0 Sv/h;

BDKG-32: 20.0 nSv/h – 0.5 Sv/h;

BDKR-01: 0.05 mSv/h – 100 mSv/h;

BDPS-02:  0.10 Sv/h – 30 mSv/h

Flux density measurement limits of α particles:

BDPA-01:  0.1 – 105 min-1 cm-2;

BDPA-02:  0.05 – 5104 min-1cm -2;

BDPS-02:  2.4 – 106 min-1cm-2;

Flux density measurement limits of β particles:

BDPB-01:  1 –5105 min-1cm-2;

BDPB-02:  0.5 – 1.5105 min-1cm-2;

BDPS-02:  6.0 – 106 min-1cm-2;

Neutron flux density measurement limits:

BDKN-01: 0.1 – 104 s-1cm-2 ;

Neutron radiation dose rate measurement limits:

BDKN-03: 0.1 mSv/h – 10.0 mSv/h;

X-ray and gamma radiation air kerma rate measurement limits:

BDKG-30: 0.02 Gy/h – 1 mGy/h;

03: ± 20 %БД
Measurement error:
PU: ± 20 %;

PU: ± 20 %;

PU-2: ± 20 %;

PU-2: ± 20 %;

PU-4: ± 20 %;

BDKG-01: ± 15 %;

BDKG-03: ± 20 %;

BDKG-04: ± 15 %;

BDKG-05: ± 15 %;

BDKG-11: ± 15 %;

BDKG-17: ± 20 %;

BDKG-24: ± 15 %;

BDKG-30: ± 15 %;

BDKG-32: ± 15 %;

BDKN-01: ± 20 %;

BDKN-03: ± 20 %;

BDPA-01: ± 20 %;

BDPA-02: ± 20 %;

BDPB-01: ± 20 %;

BDPB-02: ± 20 %;

BDKR-01: ± 15 %;

BDPS-02:  2,4 – 30 min-1 sm- 2 – ± 30 % /  30 – 106 min-1 sm- 2 – ± 20 % /  6 – 106 min-1 sm- 2 – ± 20 % /  0.10 Sv/h – 30 mSv/h – ± 20 %

Operating conditions:
- Air temperature	+$temp °C
- Atmospheric pressure	98,5 kPa
- Relative humidity	$vlag %
- Gamma radiation background	$bground nSv/h

Calibration means:
- standard dosimetry facility AT-130, N 013, Cs-137, error 5%, the Certificate of Compliance N210/1426-2018 on 22.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia);

- standard dosimetry facility AT-130, N 015, Cs-137-137, error 4 %, the Certificate of Compliance N210/ 1427-2018 on 23.10.2018 issued by FGUP "D.I.Mendeleyev VNIIM", St. Petersburg, Russia);

- collimated beams of neutrons on the standard facility UPN-AT140 №001 with the set of Pu -  - Be neutron radiation sources (certificates of Compliance N210-1508/18 issued by VNIIM, St. Petersburg, Russia), error of ± 5 %

- alpha radiation source of Pu-239: 4П9, 5П9, 6П9, error of ± 6 %, the certificateN МН0105687-4818 issued by BelGIM, Minsk, the Republic of Belarus;

- beta radiation source of Sr-90+Y-90: 4С0, 5С0, 6С0, error of 6%, the certificate N МН0100818-4818 issued by BelGIM, Minsk, the Republic of Belarus;

Calibration data

PU ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
20 Sv/h	$15-30	$15-40	$15-50	$15-60	$15-70	$15-80	$15-00	$15-10	$15-20	±15
70 Sv/h	$15-31	$15-41	$15-51	$15-61	$15-71	$15-81	$15-01	$15-11	$15-21
0,7 mSv/h	$15-32	$15-42	$15-52	$15-62	$15-72	$15-82	$15-02	$15-12	$15-22
7 mSv/h	$15-33	$15-43	$15-53	$15-63	$15-73	$15-83	$15-03	$15-13	$15-23

PU-2 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
20 Sv/h	$16-30	$16-40	$16-50	$16-60	$16-70	$16-80	$16-00	$16-10	$16-20	±15
70 Sv/h	$16-31	$16-41	$16-51	$16-61	$16-71	$16-81	$16-01	$16-11	$16-21
0,7 mSv/h	$16-32	$16-42	$16-52	$16-62	$16-72	$16-82	$16-02	$16-12	$16-22
7 mSv/h	$16-33	$16-43	$16-53	$16-63	$16-73	$16-83	$16-03	$16-13	$16-23

PU ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
20 Sv/h	$22-30	$22-40	$22-50	$22-60	$22-70	$22-80	$22-00	$22-10	$22-20	±15
70 Sv/h	$22-31	$22-41	$22-51	$22-61	$22-71	$22-81	$22-01	$22-11	$22-21
0,7 mSv/h	$22-32	$22-42	$22-52	$22-62	$22-72	$22-82	$22-02	$22-12	$22-22
7 mSv/h	$22-33	$22-43	$22-53	$22-63	$22-73	$22-83	$22-03	$22-13	$22-23
70 mSv/h	$22-34	$22-44	$22-54	$22-64	$22-74	$22-84	$22-04	$22-14	$22-24

PU-2 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
20 Sv/h	$23-30	$23-40	$23-50	$23-60	$23-70	$23-80	$23-00	$23-10	$23-20	±15
70 Sv/h	$23-31	$23-41	$23-51	$23-61	$23-71	$23-81	$23-01	$23-11	$23-21
0,7 mSv/h	$23-32	$23-42	$23-52	$23-62	$23-72	$23-82	$23-02	$23-12	$23-22
7 mSv/h	$23-33	$23-43	$23-53	$23-63	$23-73	$23-83	$23-03	$23-13	$23-23
70 mSv/h	$23-34	$23-44	$23-54	$23-64	$23-74	$23-84	$23-04	$23-14	$23-24

PU-4 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
70 Sv/h	$17-30	$17-40	$17-50	$17-60	$17-70	$17-80	$17-00	$17-10	$17-20	±15
0,7 mSv/h	$17-31	$17-41	$17-51	$17-61	$17-71	$17-81	$17-01	$17-11	$17-21
7 mSv/h	$17-32	$17-42	$17-52	$17-62	$17-72	$17-82	$17-02	$17-12	$17-22
70 mSv/h	$17-33	$17-43	$17-53	$17-63	$17-73	$17-83	$17-03	$17-13	$17-23

BDKG-01 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,7 Sv/h	$0-30	$0-40	$0-50	$0-60	$0-70	$0-80	$0-00	$0-10	$0-20	±15
7 Sv/h	$0-31	$0-41	$0-51	$0-61	$0-71	$0-81	$0-01	$0-11	$0-21
70 Sv/h	$0-32	$0-42	$0-52	$0-62	$0-72	$0-82	$0-02	$0-12	$0-22
0.7 mSv/h	$0-33	$0-43	$0-53	$0-63	$0-73	$0-83	$0-03	$0-13	$0-23
7,0 mSv/h	$0-34	$0-44	$0-54	$0-64	$0-74	$0-84	$0-04	$0-14	$0-24
70 mSv/h	$0-35	$0-45	$0-55	$0-65	$0-75	$0-85	$0-05	$0-15	$0-25
0,7 Sv/h	$0-36	$0-46	$0-56	$0-66	$0-76	$0-86	$0-06	$0-16	$0-26
7 Sv/h	$0-37	$0-47	$0-57	$0-67	$0-77	$0-87	$0-07	$0-17	$0-27

BDKG-03 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 Sv/h	$1-30	$1-40	$1-50	$1-60	$1-70	$1-80	$1-00	$1-10	$1-20	±20
0,7 Sv/h	$1-31	$1-41	$1-51	$1-61	$1-71	$1-81	$1-01	$1-11	$1-21
7 Sv/h	$1-32	$1-42	$1-52	$1-62	$1-72	$1-82	$1-02	$1-12	$1-22
70 Sv/h	$1-33	$1-43	$1-53	$1-63	$1-73	$1-83	$1-03	$1-13	$1-23
240 Sv/h	$1-34	$1-44	$1-54	$1-64	$1-74	$1-84	$1-04	$1-14	$1-24

BDKG-04 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 Sv/h	$2-30	$2-40	$2-50	$2-60	$2-70	$2-80	$2-00	$2-10	$2-20	±15
0,7 Sv/h	$2-31	$2-41	$2-51	$2-61	$2-71	$2-81	$2-01	$2-11	$2-21
7 Sv/h	$2-32	$2-42	$2-52	$2-62	$2-72	$2-82	$2-02	$2-12	$2-22
70 Sv/h	$2-33	$2-43	$2-53	$2-63	$2-73	$2-83	$2-03	$2-13	$2-23
0.7 mSv/h	$2-34	$2-44	$2-54	$2-64	$2-74	$2-84	$2-04	$2-14	$2-24
7,0 mSv/h	$2-35	$2-45	$2-55	$2-65	$2-75	$2-85	$2-05	$2-15	$2-25
70 mSv/h	$2-36	$2-46	$2-56	$2-66	$2-76	$2-86	$2-06	$2-16	$2-26
0,7 Sv/h	$2-37	$2-47	$2-57	$2-67	$2-77	$2-87	$2-07	$2-17	$2-27
7 Sv/h	$2-38	$2-48	$2-58	$2-68	$2-78	$2-88	$2-08	$2-18	$2-28

BDKG-05 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 Sv/h	$3-30	$3-40	$3-50	$3-60	$3-70	$3-80	$3-00	$3-10	$3-20	±15
0,7 Sv/h	$3-31	$3-41	$3-51	$3-61	$3-71	$3-81	$3-01	$3-11	$3-21
7 Sv/h	$3-32	$3-42	$3-52	$3-62	$3-72	$3-82	$3-02	$3-12	$3-22
70 Sv/h	$3-33	$3-43	$3-53	$3-63	$3-73	$3-83	$3-03	$3-13	$3-23
240 Sv/h	$3-34	$3-44	$3-54	$3-64	$3-74	$3-84	$3-04	$3-14	$3-24

BDKG-11 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 Sv/h	$4-30	$4-40	$4-50	$4-60	$4-70	$4-80	$4-00	$4-10	$4-20	±15
0,7 Sv/h	$4-31	$4-41	$4-51	$4-61	$4-71	$4-81	$4-01	$4-11	$4-21
7 Sv/h	$4-32	$4-42	$4-52	$4-62	$4-72	$4-82	$4-02	$4-12	$4-22
70 Sv/h	$4-33	$4-43	$4-53	$4-63	$4-73	$4-83	$4-03	$4-13	$4-23

BDKG-17 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
7,0 mSv/h	$5-30	$5-40	$5-50	$5-60	$5-70	$5-80	$5-00	$5-10	$5-20	±20
70 mSv/h	$5-31	$5-41	$5-51	$5-61	$5-71	$5-81	$5-01	$5-11	$5-21
0,7 Sv/h	$5-32	$5-42	$5-52	$5-62	$5-72	$5-82	$5-02	$5-12	$5-22
7 Sv/h	$5-33	$5-43	$5-53	$5-63	$5-73	$5-83	$5-03	$5-13	$5-23
40 Sv/h	$5-34	$5-44	$5-54	$5-64	$5-74	$5-84	$5-04	$5-14	$5-24

BDKG-24 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 Sv/h	$6-30	$6-40	$6-50	$6-60	$6-70	$6-80	$6-00	$6-10	$6-20	±15
0,7 Sv/h	$6-31	$6-41	$6-51	$6-61	$6-71	$6-81	$6-01	$6-11	$6-21
7 Sv/h	$6-32	$6-42	$6-52	$6-62	$6-72	$6-82	$6-02	$6-12	$6-22
70 Sv/h	$6-33	$6-43	$6-53	$6-63	$6-73	$6-83	$6-03	$6-13	$6-23
0.7 mSv/h	$6-34	$6-44	$6-54	$6-64	$6-74	$6-84	$6-04	$6-14	$6-24
7,0 mSv/h	$6-35	$6-45	$6-55	$6-65	$6-75	$6-85	$6-05	$6-15	$6-25
70 mSv/h	$6-36	$6-46	$6-56	$6-66	$6-76	$6-86	$6-06	$6-16	$6-26
0,7 Sv/h	$6-37	$6-47	$6-57	$6-67	$6-77	$6-87	$6-07	$6-17	$6-27

BDKG-30 ()
Kerma rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Kerma rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,03 Gy/h	$7-30	$7-40	$7-50	$7-60	$7-70	$7-80	$7-00	$7-10	$7-20	±15
0,7 Gy/h	$7-31	$7-41	$7-51	$7-61	$7-71	$7-81	$7-01	$7-11	$7-21
7 Gy/h	$7-32	$7-42	$7-52	$7-62	$7-72	$7-82	$7-02	$7-12	$7-22
70 Gy/h	$7-33	$7-43	$7-53	$7-63	$7-73	$7-83	$7-03	$7-13	$7-23
0.7 mGy/h	$7-34	$7-44	$7-54	$7-64	$7-74	$7-84	$7-04	$7-14	$7-24
7,0 mGy/h	$7-35	$7-45	$7-55	$7-65	$7-75	$7-85	$7-05	$7-15	$7-25
70 mGy/h	$7-36	$7-46	$7-56	$7-66	$7-76	$7-86	$7-06	$7-16	$7-26
0,7 Gy/h	$7-37	$7-47	$7-57	$7-67	$7-77	$7-87	$7-07	$7-17	$7-27

BDKG-32 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,07 Sv/h	$8-30	$8-40	$8-50	$8-60	$8-70	$8-80	$8-00	$8-10	$8-20	±15
0,7 Sv/h	$8-31	$8-41	$8-51	$8-61	$8-71	$8-81	$8-01	$8-11	$8-21
7 Sv/h	$8-32	$8-42	$8-52	$8-62	$8-72	$8-82	$8-02	$8-12	$8-22
70 Sv/h	$8-33	$8-43	$8-53	$8-63	$8-73	$8-83	$8-03	$8-13	$8-23
0.7 mSv/h	$8-34	$8-44	$8-54	$8-64	$8-74	$8-84	$8-04	$8-14	$8-24
7,0 mSv/h	$8-35	$8-45	$8-55	$8-65	$8-75	$8-85	$8-05	$8-15	$8-25
70 mSv/h	$8-36	$8-46	$8-56	$8-66	$8-76	$8-86	$8-06	$8-16	$8-26
0,35 Sv/h	$8-37	$8-47	$8-57	$8-67	$8-77	$8-87	$8-07	$8-17	$8-27
0,5 Sv/h	$8-38	$8-48	$8-58	$8-68	$8-78	$8-88	$8-08	$8-18	$8-28

BDKN-01 (n)
Flux density at check point 0i , s-1cm-2	Background / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,600 / 0i=5,9 / 1,0-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,600 / 0i=57,0 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,863 / 0i=235 / 2102 - 103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,957 / 0i=1019 / 2103 - 104
Readings at check point 0i, с-1см-2 , %		$9-61 / $9-71 / $9-81	$9-62 / $9-72 / $9-82	$9-63 / $9-73 / $9-83	$9-64 / $9-74 / $9-84
Average readings, i, s-1cm-2	$9-00	$9-01	$9-02	$9-03	$9-04
Result of flux density measurement at control point / npi, s-1cm-2		$9-n21	$9-n22	$9-n23	$9-n24
Intrinsic relative error of measurement, θпрi, %		$9-n31	$9-n32	$9-n33	$9-n34
Confidence limit of the intrinsic relative error Δi, %		$9-n41	$9-n42	$9-n43	$9-n44
Limite of intrinsic relative error , % not above		±20

BDKN-03 (n)
Dose rate at check point Ḣoi(10)	Bacground / R=2,0 m	ИБН-8-3 / N027 / R=2,0 m / B(R)=0,866 / H*oi(10)=8,3 / 1-10	ИБН-8-6 / N010 / R=2,0 m / B(R)=0,866 / H*oi(10)=80 / 20-100	ИБН-8-6 / N010 / R=1,0 m / B(R)=0,957 / H*oi(10)=331 / 2·102-103	ИБН-8-6 / N010 / R=0,5 m / B(R)=0,991 / H*oi(10)=1434 / 2·103-104
Readings at check point / H*i(10) , ±δст, %		$10-61 / $10-71 / $10-81	$10-62 / $10-72 / $10-82	$10-63 / $10-73 / $10-83	$10-64 / $10-74 / $10-84
Average readings, H*i(10)	$10-00	$10-01	$10-02	$10-03	$10-04
Dose rate measurement result at control point / H*прi(10)		$10-n21	$10-n22	$10-n23	$10-n24
Intrinsic relative error of measurement, θпрi, %		$10-n31	$10-n32	$10-n33	$10-n34
Confidence limit of the intrinsic relative error Δi, %		$10-n41	$10-n42	$10-n43	$10-n44
Limite of intrinsic relative error , % not above		±20

BDPA-01 (a)
Flux density in test point 0i, / min-1·cm-2	Radia-tion source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$11-90	1179	$s_bg	$11-60	$11-70	$11-80	$11-00	$11-10	$11-20	±20
$11-91	1180	---	$11-61	$11-71	$11-81	$11-01	$11-11	$11-21
$11-92	1181	---	$11-62	$11-72	$11-82	$11-02	$11-12	$11-22
$11-93	1182	---	$11-63	$11-73	$11-83	$11-03	$11-13	$11-23

BDPA-02 (a)
Flux density in test point 0i, / min-1·cm-2	Radia-tion source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$12-90	0651	$s_bg	$12-60	$12-70	$12-80	$12-00	$12-10	$12-20	±20
$12-91	0652	---	$12-61	$12-71	$12-81	$12-01	$12-11	$12-21
$12-92	0653	---	$12-62	$12-72	$12-82	$12-02	$12-12	$12-22
$12-93	0655	---	$12-63	$12-73	$12-83	$12-03	$12-13	$12-23

BDPB-01 (b)
Flux density in test point 0i, / min-1·cm-2	Radia-tion source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$13-90	0660	$b_bg	$13-60	$13-70	$13-80	$13-00	$13-10	$13-20	±20
$13-91	6905	---	$13-61	$13-71	$13-81	$13-01	$13-11	$13-21
$13-92	0661	---	$13-62	$13-72	$13-82	$13-02	$13-12	$13-22
$13-93	0662	---	$13-63	$13-73	$13-83	$13-03	$13-13	$13-23
$13-94	16437	---	$13-64	$13-74	$13-84	$13-04	$13-14	$13-24
$13-95	15014	---	$13-65	$13-75	$13-85	$13-05	$13-15	$13-25

BDPB-02 (b)
Flux density in test point 0i, / min-1·cm-2	Radia-tion source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$14-90	0663	$b2_bg	$14-60	$14-70	$14-80	$14-00	$14-10	$14-20	±20
$14-91	0664	---	$14-61	$14-71	$14-81	$14-01	$14-11	$14-21
$14-92	0665	---	$14-62	$14-72	$14-82	$14-02	$14-12	$14-22
$14-93	0661	---	$14-63	$14-73	$14-83	$14-03	$14-13	$14-23
$14-94	0666	---	$14-64	$14-74	$14-84	$14-04	$14-14	$14-24

BDKR-01 ()
Dose rate at check point Ḣoi(0.07)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(0.07)			Average value, Hi(0.07)
				H1	H2	H3
0,07 Sv/h	$18-30	$18-40	$18-50	$18-60	$18-70	$18-80	$18-00	$18-10	$18-20	±15
0,7 Sv/h	$18-31	$18-41	$18-51	$18-61	$18-71	$18-81	$18-01	$18-11	$18-21
7 Sv/h	$18-32	$18-42	$18-52	$18-62	$18-72	$18-82	$18-02	$18-12	$18-22
70 Sv/h	$18-33	$18-43	$18-53	$18-63	$18-73	$18-83	$18-03	$18-13	$18-23

BDPS-02 (a)
Flux density in test point 0i, / min-1·cm-2	Radia-tion source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$19-90	1179	$19-50	$19-60	$19-70	$19-80	$19-00	$19-10	$19-20	±20
$19-91	1180	$19-51	$19-61	$19-71	$19-81	$19-01	$19-11	$19-21
$19-92	1181	$19-52	$19-62	$19-72	$19-82	$19-02	$19-12	$19-22
$19-93	1182	$19-53	$19-63	$19-73	$19-83	$19-03	$19-13	$19-23
$19-94	1183	$19-54	$19-64	$19-74	$19-84	$19-04	$19-14	$19-24

BDPS-02 (b)
Flux density in test point 0i, / min-1·cm-2	Radia-tion source number	Measured radiation flux density in test point, i					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
		Back-ground, 0i, / min-1·cm-2	Measured value i			Average value, 
			1	2	3
$20-90	0660	$20-50	$20-60	$20-70	$20-80	$20-00	$20-10	$20-20	±20
$20-91	6905	$20-51	$20-61	$20-71	$20-81	$20-01	$20-11	$20-21
$20-92	0661	$20-52	$20-62	$20-72	$20-82	$20-02	$20-12	$20-22
$20-93	0662	$20-53	$20-63	$20-73	$20-83	$20-03	$20-13	$20-23
$20-94	15014	$20-54	$20-64	$20-74	$20-84	$20-04	$20-14	$20-24

BDPS-02 ()
Dose rate at check point Ḣoi(10)	Radia-tion source number	Distan-ce to source, / R, сm	Dose rate measurement at check point,					Relative gamma radiation dose rate measure-ment error / θпрi, %	Confidence limit of the intrinsic relative error Δi, % during calibration	Limits of intrinsic relative error, / % / not above
			Back-ground, nSv/h	Measured value Ḣ*i(10)			Average value, Hi(10)
				H1	H2	H3
0,7 Sv/h	$21-30	$21-40	$21-50	$21-60	$21-70	$21-80	$21-00	$21-10	$21-20	±20
7 Sv/h	$21-31	$21-41	$21-51	$21-61	$21-71	$21-81	$21-01	$21-11	$21-21
70 Sv/h	$21-32	$21-42	$21-52	$21-62	$21-72	$21-82	$21-02	$21-12	$21-22
0.7 mSv/h	$21-33	$21-43	$21-53	$21-63	$21-73	$21-83	$21-03	$21-13	$21-23
7,0 mSv/h	$21-34	$21-44	$21-54	$21-64	$21-74	$21-84	$21-04	$21-14	$21-24
20 mSv/h	$21-35	$21-45	$21-55	$21-65	$21-75	$21-85	$21-05	$21-15	$21-25

Calibrated by:	$calibratedBy
		(signature)
Technical control:	$controlledBy
		(signature)